ESRF Laser Drill

Safety Instructions

Instructions for Use

Maintenance

ESRF

Peter van der Linden

09/12/2009

Table of contents

- Safety instructions

  - Characteristics of the laser system

  - Safety instructions during drilling

  - Safety instructions during alignment and maintenance

2. Overview of the system

3. Instructions for use

- Maintenance

4.1 general remarks

4.2 Optical path

4.3 Alignment of the laser focal plane to the visualisation focal plane

1. Safety instructions

1.1 Characteristics of the laser system

The laser system is a class 1 system under normal use with the hood completely closed, but it becomes a class 4 sytem for maintenance and adjustment.

The characteristics of the laser are give below. This data can be useful for the application itself as well as for calculations concerning the safety of the system.

Laser head: Spectra Physics BL6S-106Q

Laser: Fiber coupled, diode pumped Q-switched Nd:YVO4

Wavelength : 1064 nm.

Power supply: Spectra Physics J40

Pulse duration: see graph below, between 6 and 14 ns pour for repetition rates between 5 kHz and 75 kHz

Pulse energy: see graph below, 100 to 200 µJ per pulse

Beam power: see graph below, between 2W and 8 W

Spatial characteristics of the beam :

Mode: TEM00

M2 <1.2

Polarisation: 100:1 vertical

Size of the beam: 0.3 mm

Divergence: 4.5 mRad

|  | Fig. 1.1: Laser power and pulse duration as a function of the pulse repetition rate |
| --- | --- |

|  | Fig. 1.2: Pulse energy as a function of the repetition rate |
| --- | --- |

The Spectra Physics model J40-BL6-106Q is a class 4 laser.

Class 4 lasers present a risk of irreversible effects on the eye and skin, and a fire risk.

Always take safety measures (in accordance with the safety instructions) to avoid exposure to direct beam and specular or diffure reflections. The fact that the infrared beam wavelength is invisible the makes it especially dangerous. The infrared radiation is focused by the lens of the eye on the retina. This can instantaneously cause internal injuries.

1.2 Safety instructions during drilling

During normal drilling operation the system is a class 1 system, no laser radiation is accessible for the user.

In order to reach this goal, the complete system is covered with a protection hood. One of the interlock contacts of the laser power supply is connected to a switch on the moving part of the hood. The laser can only be started once the moving part of the hood has been closed. The state of the laser is indicated by the two lights LASER ON and LASER OFF.

Respect the laser warning lights

When the laser is switched on, never force the hood open.

Respect the drilling procedure as described in chapter 3

Le Drilling of Be or Be containing compounds is ABSOLUTELY forbidden

Never dismount any of the fixed parts of the hood.

A marking on each panel signals the risk in case of opening :

«  Attention – class 4 laser radiation when open.

Exposure of the eye or skin to direct or diffuse radiation is dangerous.»

1.3 Safety instructions during maintenance and alignement

For all maintenance and alignment procedures the system has to be brought to a laboratory equipped with a laser safety and protection system. The laser interlock will be disconnected from the hood, and connected to the safety system of the laboratory. The laser can be started with the moving part of the hood open. The state of the laser is indicated by the two lights LASER ON and LASER OFF.

Respect the laser warning lights

Protective eyewear class L6R at 6064nm must be worn at all times during operations of maintenence and alignment.

Maintenance and alignment procedures can only be carried out by qualified personnel.

Attention! The use of controls, settings and procedures different from described in this manual can result in dangerous exposure to laser radiation.

2. Overview of the system.

Laser light path: red; illumination: yellow; visualisation: blue

3. Instructions for use

Screenshot logbook

Gasket clamp stage

Correct placement of gasket clamp

Switch on the laser power supply

Switch on the two LASCAR power supplies

Switch on the Linos piezo drive and set the piezo voltages to 70 Volt.

Switch on the monitor screen

Switch on the computer and its screen..

On the computer screen, start the logbook EXCEL file.

Start the laser control program (port1; 9600 baud)

If it gives an error try once more

Clamp the gasket in the gasket clamp stage

Align the indent to the crosshair mark.

The photo shows the XY diplacement screws

Open the hood and put the gasket clamp in place, Take care to align correctly.

A rotation by 180 degrees will also make it stick in place, but it is not correctly aligned and unstable.

Zoom button

Gasket in place

Centered and zoom to indent

Piezo adjustment

Zoom out to maximum on the screen.

The FOCUS button has no effect.

The speed button next to the zoom button allows to adjust the speed of the zoom change, normally there is no need to change it.

Zoom out and align to indent.

Start the rotation (10 -15 Volt)

Displace XY of the motor stage to center the rotation to the crosshair marks.

Stop, zoom in and adjust focus

Repeat centering until satisfied.

Push the button “p.set” on the Mitutoyo to set zero on displacement measurement

Zoom out and displace XY of gasket clamp stage to center the indent to the crosshair mark.

Zoom is chosen to fill the screen with the size of the indent.

Start rotation.

The indent will wobble on the screen.

Adjust piezo X voltage to minimuse wobble

Adjust piezo Y voltage to minimise wobble

Repeat as neccesary

At optimum a double frequency wobble appears

Not all diamonds are perfectly round!

There will always be some wobble

Rotating and offset

Drilling

Drilled

Cleaned for dimension check

Set the laser settings for the gasket material

Set the rotation to the desired voltage

Calculate:

Offset = (desired hole size)/2 - 10µm

Set the offset calculated using the Mitutoyo

Close the hood of the drill

Illumination voltage < 12V: no backlight

Illumination voiltage >12V: backlight ON

The gasket is aligned and offset

The hood is closed

Laser front panel “READY”

On the laser control start the laser

There is a three second delay

The laser starts.

When the center piece drops out the light spot is smaller, drilling is finished.

When the Backlight is ON the hole can be seen very well

Stop the laser

Stop the rotation

Open the hood

Eventually flip the gasket to check the backside

If unhappy, counterdrill

Experience is that misaligned holes mean the diamonds are misaligned

Primary gasket cleaning with alcohol has to be followed by more thorough cleaning in an ultrasound bath

4. Instructions for maintenance and alignment

4.1 general remarks

The maintenance of the different components of the system is to be carried out following the instructions as supplied by the manufacturers. This is especially valid for the laser head and power supply. All maintenance and alignment has to be carried out in a laboratory equipped with a laser safety interlock system. The interlock of the laser will be disconnected from the moving part of the protection hood, and connected to the safety system of the laboratory. The laser emission can be turned ON with the moving part of the protection hood open. The state of the laser is indicated by the two lights LASER ON and LASER OFF.

Attention! The use of controls, settings or procedures which are different from the description in this manual may result in exposition to dangerous laser radiation.

4.2 Optical path

Scheme of the optical paths

4.3 Alignment of the laser focal plane to the visualisation focal plane

The purpose of the procedure is to align the laser focal plane to the visualisation focal plane. This is done by first finding the laser focal plane and then moving the visualisation lens on the optical axis to get a sharp image.

Place a standard stainless steel gasket.

Set the laser current to 15A, frequence 15 kHz.

Find the best focus for the image on screen

Move the motor stage forward by 250 micron

Switch on the laser for a few seconds and switch off again

Move the motor stage backward by 50 micron and up by 100 micron

Switch on the laser for a few seconds and switch off again

Repeat these two steps over 11 points, ending at 250 micron behind the visualisation focus

This will create a pattern like the photo below. Values are given as examples.

If the smallest hole is drilled at the point where the image is sharp the alignment is OK.

In the example, the sharp image is at 0 µm while the smallest hole is drilled at +50 µm. This means the laser focus is at 50 µm and adjustment has to be made.

Set the motor stage to the laser focal point of +50 µm

Adjust the visualisation lens along the optical axis to get a sharp image on the monitor

Check the adjustmment between visual center and laser spot. This moves when adjusting the visualisation lens.

Compensate a horizontal deplacement by a horizontal movement of the visualisation mirror.

Compensate a vertical movement by a vertical movement of the zoom.

4.4 Alignment of the visualisation axis to the laser axis

Even if the visual and laser focalise at the same spot in the focal plane, the two axes may not be overlapping completely. This can be seen using the same initial procedure as above, without the vertical displacement. Drilling three holes at very low laser power at the focus,  200 micron in front and behind allows to evaluate if the axes are well aligned.

The optical axis of the laser has been aligned using a special infra-red visualisation aid, and it should not be adjusted.

Angular errors are to be corrected by rotation of the 45º visualisation mirror. The vertical offset error is corrected by vertical movement of the zoom; the horizontal offset error is corrected by displacement of the 45º visualisation mirror.

Perceuse Laser ESRF

Instructions de Sécurité

Instructions d’Utilisation

Maintenance

ESRF

Peter van der Linden

09/12/2009

1. Instructions de sécurité

1.1 Caractéristiques du système à laser

Système de classe 1 en utilisation normale capotages fermés ; toutefois il devient de classe 4 en maintenance ou réglages

Il utilise un laser dont les caractéristiques figurent ci-après: Ces données peuvent être utiles pour l’application mais aussi pour d’éventuels calculs liés à la sécurité.

Tête Laser: Spectra Physics BL6S-106Q

Laser: Fiber coupled, diode pumped Q-switched Nd:YVO4

Longueur d’onde 1064 nm.

Alimentation: Spectra Physics J40

Durée d’impulsion: voir graphique ci-dessous, entre 6 et 14 ns pour des taux de répétition entre 5 kHz et 75 kHz

Energie des pulses: voir graphique ci-dessous, 100 to 200 µJ par pulse

Puissance : voir graphique ci-dessous, entre 2W et 8 W

Caractéristiques spatiales du faisceau :

Mode: TEM00

M2 <1.2

Polarisation: 100:1 verticale

Taille du faisceau: 0.3 mm

Divergence: 4.5 mRad

|  | Fig. 1.1: Puissance du laser et durée des pulses en fonction du taux de répétition |
| --- | --- |

|  | Fig. 1.2: Energie par pulse en fonction du taux de répétition |
| --- | --- |

Le laser Spectra Physics modèle J40-BL6-106Q est un laser de classe 4.

Les lasers de classe 4 présentent un risque d’effets irréversibles sur l’œil, la peau et un risque d’incendie. Des mesures de sécurité doivent toujours être prises (cf instructions de sécurité) pour éviter l’exposition au faisceau direct, réfléchi spéculairement et diffusé.

Le fait que le faisceau infrarouge soit invisible le rend particulièrement dangereux. Ce rayonnement infrarouge est transmis jusqu’à la rétine en se focalisant sur celle-ci. Ceci peut instantanément causer des lésions irréversibles.

1.2 Instructions de sécurité pendant le perçage

Le système est de classe 1 en utilisation normale, en effet aucun rayonnement laser n’est accessible pour l’utilisateur.

Pour ce faire, le système laser est entièrement capoté. Un contact de sécurité du laser est relié au contacteur sur la partie mobile du capot de protection. Le laser ne peut démarrer  qu’après la fermeture du capot. L’état du laser est indiqué par les deux signaux lumineux LASER ON et LASER OFF.

Respecter les signaux d’avertissement laser

Pendant l’émission du laser ne jamais forcer l’ouverture du capot

Respecter la procédure de perçage décrite dans le chapitre3

Le perçage de Be ou alliages qui contiennent du Be est ABSOLUTEMENT interdit.

Ne jamais démonter les parties fixes du capot.

Un étiquetage mentionne le risque présenté en cas d’ouverture

«  Attention – Rayonnement laser de classe 4 en cas d’ouverture.

Exposition dangereuse de l’œil ou de la peau

Au rayonnement direct ou diffusé »

1.3 Instructions de sécurité: maintenance et alignement

Pour les opérations de maintenance et alignement le système doit être placé dans un laboratoire équipé d’un système de protection laser. Le contact de sécurité du laser sera débranché du capot, et connecté sur le système de sécurité du laboratoire. Ainsi le laser peut démarrer avec la partie mobile du capot ouvert. L’état du laser est indiqué par les deux signaux lumineux LASER ON et LASER OFF.

Respecter les signaux d’avertissement laser

Pendant le opérations de maintenance et alignement, des lunettes de protection L6R à 1064 nm sont nécessaires

- Les opérations de maintenance et alignement sont à faire uniquement par du personnel qualifié.

Attention! L’utilisation de commandes, réglages ou procédures autres que décrites dans ce manuel d’utilisation peuvent résulter en une exposition au rayonnement laser dangereux.

2. Présentation générale du système.

Chemin optique laser: rouge; illumination: jaune; visualisation: blue

- Instructions d’utilisation

Vue d’écran logbook

Platine porte joint

Placement correct

Démarrer l’alimentation du laser

Démarrer les deux alimentations LASCAR Démarrer le controleur pieze Linos piezo et ajuster les tensions des piezos à 70 Volt

Démarrer le moniteur

Démarrer l’ordinateur et son écran

Sur l’écran de l’ordinateur, démarrer le fichier EXCEL logbook

Démarrer le logiciel JWIN1 (port1; 9600 baud)

Si erreur, redemarrer encore une fois

Mettre en place le joint dans la platine porte joint

Aligner l’indentation au centre avec la croix

La photo montre les vis de déplacement XY

Ouvrir le capot et mettre en place la platine porte joint, faire attention à l’alignement

Avec une rotation de 180 degrees la platine reste en place, mais elle est n’set pas stable mail aligné

Bouton zoom

Joint placé

Centré et zoom sur indentation

Ajustement platine piezo XY

Zoom out to maximum on the screen.

The FOCUS button has no effect.

The speed button next to the zoom button allows to adjust the speed of the zoom change, normally there is no need to change it.

Faire un zoom à agrandissement minimale

Démarrer la rotation (10 -15 Volt)

Déplacer XY de la platine moteur pour avoir le centre de rotation sur la croix de moniteur.

Arreter, agrandir l’image et ajuster le focus

Répéter le centrage pour optimiser

Appuyer le bouton “p.set” du Mitutoyo pour mettre le déplacement à zéro

Reduire l’agrandissment

Déplacer XY de la platine porte joint, centre l’indentation sur la croix du moniteur

Choisir un tel agrandissement que tout l’indentation est visible sur l’écran

Démarrer la rotation.

Il y aura une oscillation de l’indentation

Minimiser l’oscillation à l’aide de tension piezo X

Minimiser l’oscillation à l’aide de tension piezo Y Répéter l’ajustement pour optimiser

Une oscillation de fréquence double apparaît autour de la réglage optimale

Les diamants ne sont pas perferctement ronds!

Il y aura toujours une petite oscillation

En rotation avec décalage

Pendant le perçage

Percé

Nettoyé pour contrôle de dimension

Dans le logiciel de contrôle laser: choisir les paramètres adaptés au matériau du joint

Ajuster la tension rotation à la valeur chosi

Calcul de décalage:

Décalage = (taille du trou)/2 - 10µm

Déplacer la platine moteur pour faire le décalage à l’aide du Mitutoyo

Fermer le capot

Le joint a été aligné et décalé

Le capot est fermé

La face avant de l’alimentation laser: “READY”

Démarrer le laser avec le logiciel de contrôle

Il y a un délai de trois secondes

Le laser démarre

Tension d’illumination <12V: lumière arrière eteinte

Tension d’illumination >12V : lumière arrière allumé

Quand la partie centrale tombe et la tâche lumineuse reduit de taille le precage est fini

Avec la lumière arrière allumé le trou devient tres bien visible

Arrêter le laser

Ouvrir le capot

Eventuellement retourner le joint pour un controle visuel de l’arrière

Si le résultat ne convient pas, repercer

L’expérience est que des trous mal alignés sont causés par des diamants mal alignés

Apres le premier nettoyage sur place il faut faire un nettoyage dans un bain d’ultrason

4. Instructions d’entretien et alignement

4.1 Remarques générales

L’entretien des différents composants du système est à faire selon les préconisations et intructions des différents fabricants. Ceci est espécialement important en ce qui concerne la tête et l’alimentation du laser. Toutes les opérations d’entretien et alignement sont à faire dans un laboratoire équipé d’un système de sécurité laser. Le contact de sécurité du laser sera débranché du capot, et connecté sur le système de sécurité du laboratoire. Ainsi le laser peut démarrer avec la partie mobile du capot ouvert. L’état du laser est indiqué par les deux signes lumineuses LASER ON et LASER OFF.

Attention! L’utilisation de commandes, réglages ou procedures autres que décrites dans ce manuel d’utilisation peuvent resulter en une exposition au rayonnement laser dangereux

4.2 Chemins optiques

Schematique des chemins optiques

4.3 Alignment du plan focal du laser avec le plan focal de la visualisation

Le but de cette procédure est l’alignement du plan focal du laser sur le plan focal de la visualisation. La première chose à faire est de trouver le plan focal du laser; ensuite la position de la lentille de visualisation est pour avoir une iomage nette sur le moniteur.

Monter un joint standard en inox.

Ajuster le courant du lasr sur 15A, et la fréquence sur 15kHz.

Ajuster la focale de la platine moteur pour faire une image nette

Avancer la platine moteur de 200 micron sur l’axe optique avec la commande focale

Démarrer le laser quelques secondes et l’arrêter, il suffit de marquer la surface

Reculer la platine moteur de 50 micron faire un mouvement vertical de 100 micron

Démarrer le laser quelques secondes et l’arrêter

Répéter ceci en total neuf fois, pour terminer à 200  micron derrière le plan focal.

Les marques succesifs produisent un motif comme ci-dessous. Les valeurs sont donnés comme examples.

Si le plus petit trou est percé là ou l’image est nette, le réglage est bon.

Dans l’example, l’image nette se trouve à 0 µm tandis que le plus petit trou est percé à +50µm. Dans ce cas, il y a un décalage de 50 µm entre le plan focal du laser et le plan focal de la visualisation.

Ajuster la platine moteur au plan focal du laser à +50 µm

Ajuster ensuite la lentille de visualisation sur l’axe optique pour avoir une image nette

L’étape finale est la vérification que l’axe de visualisation n’a pas été déplacé.

Refaire une autre marque, laser allumé quelques secondes

Cette marque doit se faire bien centré sur la croix.

Compenser un déplacement horizontal à l’aide d’un déplacement horizontal du miroir de visualisation

Compenser un déplacement vertical à l’aide d’un déplacement vertical du zoom

4.4 Alignment of the visualisation axis to the laser axis

Even if the visual and laser focalise at the same spot in the focal plane, the two axes may not be overlapping completely. This can be seen using the same initial procedure as above, without the vertical displacement. Drilling three holes at very low laser power at the focus,  200 micron in front and behind allows to evaluate if the axes are well aligned.

The optical axis of the laser has been aligned using a special infra-red visualisation aid, and it should not be adjusted.

Angular errors are to be corrected by rotation of the 45º visualisation mirror. The vertical offset error is corrected by vertical movement of the zoom; the horizontal offset error is corrected by displacement of the 45º visualisation mirror.

# Text boxes

XY

XY

XY of

gasket clamp

stage

XY of

motor stage

FOCUS

PIEZO

Do not force

OBJECTIVE

Do not touch

Keyboard

Zoom control

Rotation

Light

Monitor

Mitutoyo

Piezo drive

Laser head

Motor stage

Focusing objective

Screen

Laser head

Illumination

Zoom and camera

Visualisation mirror

Visualisation lens

IR mirror

Beam splitter

Objective

Laser head

Condensor lens

Illumination diode

Beam expander

Gasket clamp

stage

Zoom and camera

+100μm

+50μm

Sharp image

-50μm

Visualisation light source

Condensor lens

Zoom and camera

IR mirror

Focal plane

Objective

Beam splitter

Visualisation lens

Visualisation mirror

Laser and beam expander

Three spots overlap

Alignment OK

Crosshair on monitor

Three distinct spots

Alignment wrong

Vertical error

Horizontal error

Control du zoom

Rotation

Lumière

Objectif pour focalisation

Ecran

Moniteur

Mitutoyo

Controleur Piezo

Tête laser

Moteur

Clavier

Zoom et camera

Illumination

Tête laser

Zoom et camera

Lentille Condensor

Illumination diode

Tête laser

Téléscope

Platine joint

Objectif

Miroir IR

Beam splitter

Lentille de Visualisation

Miroir de visualisation

PIEZO

Ne pas forcer

XY

Platine moteur

XY

Platine joint

OBJECTIF

Ne pas toucher

FOCUS

XY

XY

Tête laser et téléscope

Miroir de

Visualisation

Lentille de visualisation

Beam splitter

Objectif

Plan focal

Miroir IR

Zoom et camera

Condensor lens

Lumière de visualisation

-50μm

Sharp image

+50μm

+100μm

Three distinct spots

Alignment wrong

Three spots overlap

Alignment OK

Crosshair on monitor

Vertical error

Horizontal error